Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Glassware Washing and Drying - Wet Lab

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-302-025 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-302-025 Glassware Washing and Drying - Wet Lab |
| Parent procedure | HM-PRC-VXX-PRO213 Laboratory Wet Chemistry and Digestion |
| Area | 302 Wet Lab |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-302-025 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task glassware washing and drying - wet lab in 302  Wet Lab at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Wash and dry wet lab glassware.

- A trace of gold or acid carried over on glassware contaminates the next result invisibly.

It is read with work instruction SWI-PRO-LAB-302-025, which gives the step-by-step method, and with HM-PRC-VXX-PRO213 Laboratory Wet Chemistry and Digestion, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Rinse immediately after use, before residues dry on. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Cross-contamination from a container or cutter not cleaned between points | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 2. Discard the first rinse to the correct waste stream - an acid or cyanide rinse is not general drain. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Hydrogen cyanide gas released at the open sample point Inhalation of acid fume or reagent vapour at the bench | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 3 | 5 | 15 EXTREME | Laboratory Technician |
| 3. Wash with detergent and a brush, then rinse with tap water. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Cross-contamination from a container or cutter not cleaned between points | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 4. Acid wash where the method requires it, in the fume cupboard, then rinse thoroughly. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 4 | 12 HIGH | Laboratory Technician |
| 5. Give a final rinse with distilled or RO/DI water. Tap water leaves dissolved solids behind. | Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 6. Inspect for chips and cracks and discard damaged glassware into the sharps bin. | Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 7. Do NOT dry volumetric glassware in an oven - heat changes its calibrated volume. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Cross-contamination from a container or cutter not cleaned between points | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 8. Store clean glassware inverted or covered. | Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: After every use.

6.2  Job steps

1.  Rinse immediately after use, before residues dry on.

2.  Discard the first rinse to the correct waste stream - an acid or cyanide rinse is not general drain.

3.  Wash with detergent and a brush, then rinse with tap water.

4.  Acid wash where the method requires it, in the fume cupboard, then rinse thoroughly.

5.  Give a final rinse with distilled or RO/DI water. Tap water leaves dissolved solids behind.

6.  Inspect for chips and cracks and discard damaged glassware into the sharps bin.

7.  Do NOT dry volumetric glassware in an oven - heat changes its calibrated volume.

8.  Store clean glassware inverted or covered.

Equipment required:

- Wash sinks

- Detergent and brushes

- Acid wash bath

- Distilled or RO/DI water

- Drying oven and racks

- Fume cupboard

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-302-025 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-302-025.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-WET-SOP-049 drying and ENV-SOP-060 washing

- Heritage register line JSEA-PRO-LAB-302-025, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-302-025 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |