Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Perchloric Acid Digest and Fume Cupboard Washdown Routine

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-302-032 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-302-032 Perchloric Acid Digest and Fume Cupboard Washdown Routine |
| Parent procedure | HM-PRC-VXX-PRO213 Laboratory Wet Chemistry and Digestion |
| Area | 302 Wet Lab |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-302-032 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task perchloric acid digest and fume cupboard washdown routine in 302  Wet Lab at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Carry out a perchloric acid digest and the mandatory fume cupboard washdown.

- Perchloric acid fume condenses in ductwork as metal perchlorates, which are shock-sensitive explosives. The washdown is not housekeeping - it is what prevents an explosion.

It is read with work instruction SWI-PRO-LAB-302-032, which gives the step-by-step method, and with HM-PRC-VXX-PRO213 Laboratory Wet Chemistry and Digestion, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Use ONLY the fume cupboard rated and plumbed for perchloric acid. Never use a general cupboard, and never use one where organic solvent work has been done. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 2. Confirm the washdown system is functional before starting. If it is not, do not start. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 3. Never mix perchloric acid with organic material. Perchloric acid and organics together is an explosion, not a reaction. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 4. Where the sample contains organic matter, destroy it with nitric acid first and confirm destruction before adding perchloric. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 5. Take to fumes as the method directs but NEVER to dryness. A perchloric digest taken to dryness is the classic explosion. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 4 | 12 HIGH | Laboratory Technician |
| 6. Keep the digest under observation for the whole fuming stage. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 7. Run the fume cupboard washdown immediately after the digest, for the full stated period. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 4 | 12 HIGH | Laboratory Technician |
| 8. Record the digest and the washdown in the log. An unrecorded washdown is treated as not done. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 4 | 12 HIGH | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: Every perchloric digest.

6.2  Job steps

1.  Use ONLY the fume cupboard rated and plumbed for perchloric acid. Never use a general cupboard, and never use one where organic solvent work has been done.

2.  Confirm the washdown system is functional before starting. If it is not, do not start.

3.  Never mix perchloric acid with organic material. Perchloric acid and organics together is an explosion, not a reaction.

4.  Where the sample contains organic matter, destroy it with nitric acid first and confirm destruction before adding perchloric.

5.  Take to fumes as the method directs but NEVER to dryness. A perchloric digest taken to dryness is the classic explosion.

6.  Keep the digest under observation for the whole fuming stage.

7.  Run the fume cupboard washdown immediately after the digest, for the full stated period.

8.  Record the digest and the washdown in the log. An unrecorded washdown is treated as not done.

Equipment required:

- Fume cupboard rated and plumbed for perchloric acid

- Washdown system

- Hotplate

- Acid PPE and face shield

- pH paper

- Log sheet

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-302-032 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-302-032.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- Perchloric washdown is mandatory to prevent explosive perchlorate build-up in ductwork. No procedure held and no washdown provision itemised in ILS quote Q0023979

- Heritage register line JSEA-PRO-LAB-302-032, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-302-032 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |