Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Oxygen and Argon Supply, Regulator and Flashback Arrestor Check

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-303-013 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-303-013 Oxygen and Argon Supply, Regulator and Flashback Arrestor Check |
| Parent procedure | HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis |
| Area | 303 AAS and Instruments |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-303-013 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task oxygen and argon supply, regulator and flashback arrestor check in 303  AAS and Instruments at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Check the oxygen and argon supply, regulators and flashback arrestors.

- Oxygen accelerates any fire, and a contaminated oxygen fitting can ignite on its own.

It is read with work instruction SWI-PRO-LAB-303-013, which gives the step-by-step method, and with HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Confirm each cylinder is restrained and the correct regulator is fitted for the gas. Regulators are not interchangeable between gases. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 2. NEVER allow oil or grease near an oxygen fitting. Oxygen and hydrocarbon ignite on contact under pressure. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 3. Confirm every flashback arrestor is fitted, correctly oriented and within its service life. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 4. Open cylinder valves slowly, standing to the side. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 5. Leak test every joint with leak detection solution, never a flame. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 6. Record supply pressures and compare against the previous check. A falling standing pressure is a leak. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 7. Confirm cylinder change-out before a cylinder runs empty - drawing a cylinder to zero pulls contamination into the line. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 8. Record the check in the log. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: At the set interval and at every cylinder change.

6.2  Job steps

1.  Confirm each cylinder is restrained and the correct regulator is fitted for the gas. Regulators are not interchangeable between gases.

2.  NEVER allow oil or grease near an oxygen fitting. Oxygen and hydrocarbon ignite on contact under pressure.

3.  Confirm every flashback arrestor is fitted, correctly oriented and within its service life.

4.  Open cylinder valves slowly, standing to the side.

5.  Leak test every joint with leak detection solution, never a flame.

6.  Record supply pressures and compare against the previous check. A falling standing pressure is a leak.

7.  Confirm cylinder change-out before a cylinder runs empty - drawing a cylinder to zero pulls contamination into the line.

8.  Record the check in the log.

Equipment required:

- 2-stage pure oxygen and argon regulators

- Acetylene regulators

- Flashback arrestors

- Leak detection solution

- Cylinder restraints

- Log sheet

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-303-013 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-303-013.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- ILS quote Q0023979: 2-stage pure oxygen/argon regulator, acetylene regulators and flashback arrestors; KBK AAS-SOP-066 is the pattern

- Heritage register line JSEA-PRO-LAB-303-013, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-303-013 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |