Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

AAS Sample Reading - Base Metals

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-303-003 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-303-003 AAS Sample Reading - Base Metals |
| Parent procedure | HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis |
| Area | 303 AAS and Instruments |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-303-003 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task aas sample reading - base metals in 303  AAS and Instruments at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Read base metals on the AAS.

- The quote is flame AAS only - there is no graphite furnace and no ICP, so the detection limit is what flame gives.

It is read with work instruction SWI-PRO-LAB-303-003, which gives the step-by-step method, and with HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Fit the correct lamp and allow the warm-up period. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 2. Match the standard matrix to the sample matrix. A matrix mismatch shifts the whole calibration and nothing in the data shows it. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 3. Set the method parameters from the PerkinElmer method for that element. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 4. Calibrate and confirm the correlation before reading samples. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 5. Read quality control samples at the specified positions. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 6. Aspirate blank between samples. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 7. Check for interferences the method warns about - iron on copper is the common one in this circuit. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 8. Record dilutions and any interference correction applied. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: Every base metal batch.

6.2  Job steps

1.  Fit the correct lamp and allow the warm-up period.

2.  Match the standard matrix to the sample matrix. A matrix mismatch shifts the whole calibration and nothing in the data shows it.

3.  Set the method parameters from the PerkinElmer method for that element.

4.  Calibrate and confirm the correlation before reading samples.

5.  Read quality control samples at the specified positions.

6.  Aspirate blank between samples.

7.  Check for interferences the method warns about - iron on copper is the common one in this circuit.

8.  Record dilutions and any interference correction applied.

Equipment required:

- PinAAcle 500 with Syngistix

- Element hollow cathode lamps

- Standards matched to the sample matrix

- Sample racks

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-303-003 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-303-003.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- ILS quote Q0023979: PinAAcle 500, flame only - no graphite furnace or ICP quoted

- Heritage register line JSEA-PRO-LAB-303-003, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-303-003 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |