Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Acetylene and Air Cylinder Changeover and Leak Test

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-303-005 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-303-005 Acetylene and Air Cylinder Changeover and Leak Test |
| Parent procedure | HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis |
| Area | 303 AAS and Instruments |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-303-005 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task acetylene and air cylinder changeover and leak test in 303  AAS and Instruments at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Change over acetylene and air cylinders and leak test the connections.

- Acetylene is flammable across a very wide range in air and is unstable under pressure. A leaking connection in an enclosed laboratory is a serious event.

It is read with work instruction SWI-PRO-LAB-303-005, which gives the step-by-step method, and with HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Shut down the flame and close the cylinder valve before disconnecting anything. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 2. Move cylinders on the trolley with the valve cap fitted. Never roll or drag a cylinder. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 3. Restrain the cylinder before removing the cap. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 4. Confirm the regulator is the correct type for the gas and is within its service date. Never use an oxygen regulator on acetylene. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 5. Confirm the flashback arrestor is fitted and in date. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 6. Do not use oil or grease on any fitting. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 7. Open the cylinder valve slowly, standing to the side of the regulator. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 8. Leak test every joint with leak detection solution. Never with a flame. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 9. Do NOT exceed the acetylene withdrawal rate the cylinder allows - drawing too fast pulls acetone over into the line. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 10. Record the cylinder numbers and the change in the log. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: Every cylinder change.

6.2  Job steps

1.  Shut down the flame and close the cylinder valve before disconnecting anything.

2.  Move cylinders on the trolley with the valve cap fitted. Never roll or drag a cylinder.

3.  Restrain the cylinder before removing the cap.

4.  Confirm the regulator is the correct type for the gas and is within its service date. Never use an oxygen regulator on acetylene.

5.  Confirm the flashback arrestor is fitted and in date.

6.  Do not use oil or grease on any fitting.

7.  Open the cylinder valve slowly, standing to the side of the regulator.

8.  Leak test every joint with leak detection solution. Never with a flame.

9.  Do NOT exceed the acetylene withdrawal rate the cylinder allows - drawing too fast pulls acetone over into the line.

10.  Record the cylinder numbers and the change in the log.

Equipment required:

- Acetylene and air cylinders

- Cylinder trolley and restraints

- Regulators and flashback arrestors

- Leak detection solution

- Spanners - non-sparking

- Cylinder log

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-303-005 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-303-005.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-AAS-SOP-066 Penggantian tabung gas asetilen

- Heritage register line JSEA-PRO-LAB-303-005, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-303-005 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |