Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

AAS Burner and Nebuliser Cleaning and Maintenance

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-303-006 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-303-006 AAS Burner and Nebuliser Cleaning and Maintenance |
| Parent procedure | HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis |
| Area | 303 AAS and Instruments |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-303-006 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task aas burner and nebuliser cleaning and maintenance in 303  AAS and Instruments at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Clean and maintain the AAS burner and nebuliser.

- A partly blocked nebuliser reduces sensitivity gradually. The results stay plausible while they drift low.

It is read with work instruction SWI-PRO-LAB-303-006, which gives the step-by-step method, and with HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Shut down the flame and allow the burner to cool completely before touching it. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 2. Remove the burner head and inspect the slot. Carbon build-up along the slot is the usual cause of a noisy flame. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 3. Clean the slot with the cleaning card only. Never use a wire or a blade - a scratched slot is a ruined burner head. | Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 4. Remove and inspect the nebuliser. Clean it in the ultrasonic bath with distilled water. | Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 5. Replace o-rings at the interval or on any sign of hardening. A leaking o-ring gives a wandering signal. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 6. Check and refill the drain trap. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 7. Reassemble, re-ignite and confirm sensitivity against a standard before running samples. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 8. Record the service and the post-service sensitivity in the log. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: At the interval the method sets and whenever sensitivity falls.

6.2  Job steps

1.  Shut down the flame and allow the burner to cool completely before touching it.

2.  Remove the burner head and inspect the slot. Carbon build-up along the slot is the usual cause of a noisy flame.

3.  Clean the slot with the cleaning card only. Never use a wire or a blade - a scratched slot is a ruined burner head.

4.  Remove and inspect the nebuliser. Clean it in the ultrasonic bath with distilled water.

5.  Replace o-rings at the interval or on any sign of hardening. A leaking o-ring gives a wandering signal.

6.  Check and refill the drain trap.

7.  Reassemble, re-ignite and confirm sensitivity against a standard before running samples.

8.  Record the service and the post-service sensitivity in the log.

Equipment required:

- High-sensitivity nebuliser and o-ring kits

- Burner cleaning card

- Ultrasonic bath

- Distilled water

- Soft cloth

- Maintenance log

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-303-006 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-303-006.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- ILS quote Q0023979 includes high-sensitivity nebulisers and o-ring kits but no maintenance procedure is held

- Heritage register line JSEA-PRO-LAB-303-006, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-303-006 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |