Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Instrument Fault Response and Sample Re-Run Protocol

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-303-008 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-303-008 Instrument Fault Response and Sample Re-Run Protocol |
| Parent procedure | HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis |
| Area | 303 AAS and Instruments |
| Classification | Operating |
| Register line | JSEA-PRO-LAB-303-008 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task instrument fault response and sample re-run protocol in 303  AAS and Instruments at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Respond to an instrument fault and decide what has to be re-run.

- The question after a fault is not just 'is it fixed' but 'which results can still be trusted'.

It is read with work instruction SWI-PRO-LAB-303-008, which gives the step-by-step method, and with HM-PRC-VXX-PRO214 Laboratory Instrumental Analysis, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Stop the run. Do not keep reading through a fault hoping it clears. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 2. Record what the instrument was doing, what the last good quality control result was, and the time. | Contact with live electrical equipment at a sample pump, cutter or local control station Non-representative sample leading to a wrong process decision | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Laboratory Technician |
| 3. Establish the last point at which the instrument was demonstrably reading correctly - the last in-control standard or drift check. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 4. Treat every sample read after that point as suspect until proven otherwise. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 5. Diagnose against the manual. Do not open an instrument enclosure or attempt an electrical repair. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 6. After the fix, recalibrate fully and confirm with a standard and a blank before restarting. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Laboratory Technician |
| 7. Re-run every suspect sample. Do not release results from the suspect window. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 8. Record the fault, the cause, the results withdrawn and the results re-run in the fault register. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: Every instrument fault.

6.2  Job steps

1.  Stop the run. Do not keep reading through a fault hoping it clears.

2.  Record what the instrument was doing, what the last good quality control result was, and the time.

3.  Establish the last point at which the instrument was demonstrably reading correctly - the last in-control standard or drift check.

4.  Treat every sample read after that point as suspect until proven otherwise.

5.  Diagnose against the manual. Do not open an instrument enclosure or attempt an electrical repair.

6.  After the fix, recalibrate fully and confirm with a standard and a blank before restarting.

7.  Re-run every suspect sample. Do not release results from the suspect window.

8.  Record the fault, the cause, the results withdrawn and the results re-run in the fault register.

Equipment required:

- Instrument logs

- Batch sheets

- Quality control data

- Fault register

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-303-008 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-303-008.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- Heritage register line JSEA-PRO-LAB-303-008, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-303-008 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |