Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

pH Meter Calibration and Use

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-LAB-304-001 |
| --- | --- |
| Work instruction | SWI-PRO-LAB-304-001 pH Meter Calibration and Use |
| Parent procedure | HM-PRC-VXX-PRO215 Laboratory Environmental and Cyanide Analysis |
| Area | 304 Environmental Analysis |
| Classification | Safety-critical |
| Register line | JSEA-PRO-LAB-304-001 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_LAB |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task ph meter calibration and use in 304  Environmental Analysis at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Calibrate and use the pH meter.

- pH governs cyanide safety as much as it governs chemistry. An uncalibrated pH meter in a cyanide laboratory is a safety issue.

It is read with work instruction SWI-PRO-LAB-304-001, which gives the step-by-step method, and with HM-PRC-VXX-PRO215 Laboratory Environmental and Cyanide Analysis, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Check the buffer expiry dates. An out-of-date buffer calibrates the meter to the wrong value. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 2. Rinse the electrode with distilled water and blot dry between buffers. Never wipe - it strips the junction. | Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 3. Calibrate across pH 4, 7 and 10 and record the slope and offset. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Non-representative sample leading to a wrong process decision | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 4 | 12 HIGH | Laboratory Technician |
| 4. Reject the electrode where the slope falls outside the acceptable range rather than continuing because it still gives a number. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 5. Record the temperature at calibration. pH is temperature dependent. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Non-representative sample leading to a wrong process decision | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 4 | 12 HIGH | Laboratory Technician |
| 6. Rinse between samples and allow the reading to stabilise before recording. | Cross-contamination from a container or cutter not cleaned between points Non-representative sample leading to a wrong process decision | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |
| 7. Store the electrode in storage solution, never in distilled water. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Laboratory Technician |
| 8. Record every calibration in the meter log against the electrode identity. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Laboratory Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Frequency: Every shift before use.

6.2  Job steps

1.  Check the buffer expiry dates. An out-of-date buffer calibrates the meter to the wrong value.

2.  Rinse the electrode with distilled water and blot dry between buffers. Never wipe - it strips the junction.

3.  Calibrate across pH 4, 7 and 10 and record the slope and offset.

4.  Reject the electrode where the slope falls outside the acceptable range rather than continuing because it still gives a number.

5.  Record the temperature at calibration. pH is temperature dependent.

6.  Rinse between samples and allow the reading to stabilise before recording.

7.  Store the electrode in storage solution, never in distilled water.

8.  Record every calibration in the meter log against the electrode identity.

Equipment required:

- pH meter and electrode

- Buffers 4, 7 and 10, in date

- Distilled or RO/DI water

- Soft tissue

- Storage solution

- Meter log

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-LAB-304-001 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-LAB-304-001.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-ENV-SOP-067, -068 calibration and -087 reading

- Heritage register line JSEA-PRO-LAB-304-001, tab JSEA_LAB / SWI-LAB

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-LAB-304-001 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |