Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Automatic Titrator Operation and Method Set-Up

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-205-018 |
| --- | --- |
| Work instruction | SWI-PRO-MET-205-018 Automatic Titrator Operation and Method Set-Up |
| Parent procedure | HM-PRC-VXX-PRO209 Metallurgical Laboratory Chemical and Leach Testwork |
| Area | 205 Met Lab - Chemical and Leach Testwork |
| Classification | Operating |
| Register line | JSEA-PRO-MET-205-018 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task automatic titrator operation and method set-up in 205  Met Lab - Chemical and Leach Testwork at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Set up and run the Hanna HI902C automatic titrator with the methods the plant needs.

- The titrator replaces manual titration for routine work. A wrongly configured method produces confident wrong answers all shift.

It is read with work instruction SWI-PRO-MET-205-018, which gives the step-by-step method, and with HM-PRC-VXX-PRO209 Metallurgical Laboratory Chemical and Leach Testwork, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Confirm the correct electrode or ISE is fitted for the determination and that it is conditioned. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 2. Load the titrant and prime the burette, expelling all air. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 3. Enter the method parameters from the manual method being ported - aliquot, titrant concentration, endpoint criterion, stirring and dosing rate. | Hydrogen cyanide gas released at the open sample point Contact with live electrical equipment at a sample pump, cutter or local control station | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 4. Run the method against a known standard first. Do not run samples until the standard reads correctly. | Hydrogen cyanide gas released at the open sample point Non-representative sample leading to a wrong process decision | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 5. Compare the titrator result against a manual titration on the same sample, in duplicate. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Contact with sodium hydroxide solution - burn to skin or eyes Hydrogen cyanide gas released at the open sample point | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Full chemical PPE including face shield and chemical-resistant gloves. Sample drawn slowly from the designated valve. Safety shower and eyewash within reach and proven flowing. Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 6. Where the two disagree beyond tolerance, correct the method - do not adjust the result. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 7. Save the method under a controlled name and version and record it in the method register. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 8. Re-verify against a manual titration weekly and after any electrode change. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Contact with sodium hydroxide solution - burn to skin or eyes Hydrogen cyanide gas released at the open sample point | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Full chemical PPE including face shield and chemical-resistant gloves. Sample drawn slowly from the designated valve. Safety shower and eyewash within reach and proven flowing. Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: At set-up, and whenever a method is added or changed.

6.2  Job steps

1.  Confirm the correct electrode or ISE is fitted for the determination and that it is conditioned.

2.  Load the titrant and prime the burette, expelling all air.

3.  Enter the method parameters from the manual method being ported - aliquot, titrant concentration, endpoint criterion, stirring and dosing rate.

4.  Run the method against a known standard first. Do not run samples until the standard reads correctly.

5.  Compare the titrator result against a manual titration on the same sample, in duplicate.

6.  Where the two disagree beyond tolerance, correct the method - do not adjust the result.

7.  Save the method under a controlled name and version and record it in the method register.

8.  Re-verify against a manual titration weekly and after any electrode change.

Equipment required:

- Hanna HI902C pH/ORP/ISE titrator

- Electrodes and ISE probes

- Standard solutions

- Titrant reservoirs

- Beakers and stirrers

- Fume cupboard

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-205-018 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-205-018.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- ILS quotation Q0023979 - Hanna HI902C automatic titrator with ISE

- KBK-MIR-MP-PRO-MET-SOP-0025, -0037, -0039, -0040, -0041, -0042 and -0043 - the seven titration methods to port, all received 11 Aug 2026

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-205-018 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |