Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Centrifuge Operation for Solution Clarification

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-204-019 |
| --- | --- |
| Work instruction | SWI-PRO-MET-204-019 Centrifuge Operation for Solution Clarification |
| Parent procedure | HM-PRC-VXX-PRO208 Metallurgical Laboratory Physical Testwork |
| Area | 204 Met Lab - Physical Testwork |
| Classification | Operating |
| Register line | JSEA-PRO-MET-204-019 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task centrifuge operation for solution clarification in 204  Met Lab - Physical Testwork at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Clarify a solution sample by centrifuge where filtration is impractical or too slow.

- Used where a fine or gelatinous solid blinds a filter paper.

It is read with work instruction SWI-PRO-MET-204-019, which gives the step-by-step method, and with HM-PRC-VXX-PRO208 Metallurgical Laboratory Physical Testwork, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Inspect the tubes for cracks before loading. A cracked tube fails at speed. | Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 2. Fill opposing tubes to the same level and balance them on the balance to within the manufacturer's tolerance. An unbalanced rotor is the main failure mode of a centrifuge. | Contact with an agitator, sample cutter, pump or screen that starts without warning Non-representative sample leading to a wrong process decision | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 3. Load opposing positions symmetrically. Never run with a single tube. | Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 4. Close and latch the lid before starting. | Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 5. Run at the speed and for the period the method specifies. | Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 6. Allow the rotor to come to a complete stop on its own. Never brake it by hand. | Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 7. Decant the clarified supernatant without disturbing the pellet. | Splash of cyanide-bearing slurry or solution to skin or eyes | Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. | 3 | 4 | 12 HIGH | Met Lab Technician |
| 8. Record speed, time and the tube set used. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: As required by the analysis schedule.

6.2  Job steps

1.  Inspect the tubes for cracks before loading. A cracked tube fails at speed.

2.  Fill opposing tubes to the same level and balance them on the balance to within the manufacturer's tolerance. An unbalanced rotor is the main failure mode of a centrifuge.

3.  Load opposing positions symmetrically. Never run with a single tube.

4.  Close and latch the lid before starting.

5.  Run at the speed and for the period the method specifies.

6.  Allow the rotor to come to a complete stop on its own. Never brake it by hand.

7.  Decant the clarified supernatant without disturbing the pellet.

8.  Record speed, time and the tube set used.

Equipment required:

- Ohaus FC5706 centrifuge

- 15 mL centrifuge tubes

- Balance for tube pairing

- Timer

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-204-019 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-204-019.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- ILS quotation Q0023979 - Ohaus FC5706 centrifuge and 15 mL tubes

- No procedure was held for this instrument before this revision

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-204-019 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |