Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Analyser to Laboratory Correlation Check and Drift Investigation

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-207-012 |
| --- | --- |
| Work instruction | None. This is a desk-based task with a JSEA line but no work instruction on the SWI-MET register. |
| Parent procedure | HM-PRC-VXX-PRO211 Automatic Sampling and Online Analyser Operation |
| Area | 207 Automatic Sampling and Online Analysers |
| Classification | Safety-critical |
| Register line | JSEA-PRO-MET-207-012 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task analyser to laboratory correlation check and drift investigation in 207  Automatic Sampling and Online Analysers at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Track the online analyser reading against the laboratory result over time, and investigate drift before it becomes a control problem.

- An analyser that drifts slowly is more dangerous than one that fails, because the plant keeps being controlled to it.

It is read with HM-PRC-VXX-PRO211 Automatic Sampling and Online Analyser Operation, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Take the manual sample from the analyser sample line, or as close to it as the arrangement allows, and record the analyser reading at that instant. | Hydrogen cyanide gas released at the open sample point Contact with live electrical equipment at a sample pump, cutter or local control station Non-representative sample leading to a wrong process decision | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Operator |
| 2. Determine free cyanide by the laboratory method, and WAD cyanide where it can be determined. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Operator |
| 3. Plot the analyser reading against the laboratory result on a running chart. | Hydrogen cyanide gas released at the open sample point Contact with live electrical equipment at a sample pump, cutter or local control station Non-representative sample leading to a wrong process decision | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Operator |
| 4. Look at the trend, not the single point. A one-off difference is noise; a drift in one direction over weeks is a real problem. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Operator |
| 5. Where drift is developing, check the electrolyte, the electrode age and the sample line condition before requesting a recalibration. | Hydrogen cyanide gas released at the open sample point Non-representative sample leading to a wrong process decision | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Operator |
| 6. Report any analyser outside its tolerance to instrumentation, and record the interim correction the control room should apply. | Hydrogen cyanide gas released at the open sample point Contact with live electrical equipment at a sample pump, cutter or local control station Non-representative sample leading to a wrong process decision | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Operator |
| 7. Never adjust the analyser to match the laboratory without confirming the laboratory first - the laboratory can also be wrong. | Hydrogen cyanide gas released at the open sample point Contact with live electrical equipment at a sample pump, cutter or local control station | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 3 | 5 | 15 EXTREME | Operator |
| 8. Report the correlation position monthly. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Operator |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Harness and fall arrest where no fixed platform exists

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| 032-CA-001 and 051-CA-002 | Online cyanide analysers |
| Laboratory | Free and WAD cyanide by the manual methods |

Frequency: Weekly correlation; investigation on any developing drift.

6.2  Job steps

1.  Take the manual sample from the analyser sample line, or as close to it as the arrangement allows, and record the analyser reading at that instant.

2.  Determine free cyanide by the laboratory method, and WAD cyanide where it can be determined.

3.  Plot the analyser reading against the laboratory result on a running chart.

4.  Look at the trend, not the single point. A one-off difference is noise; a drift in one direction over weeks is a real problem.

5.  Where drift is developing, check the electrolyte, the electrode age and the sample line condition before requesting a recalibration.

6.  Report any analyser outside its tolerance to instrumentation, and record the interim correction the control room should apply.

7.  Never adjust the analyser to match the laboratory without confirming the laboratory first - the laboratory can also be wrong.

8.  Report the correlation position monthly.

Equipment required:

- Sample containers

- Free cyanide titration set

- Analyser verification log

- Trend charts

# 7  FLOWCHART

Not applicable. This is a desk-based task and the step sequence is in Section 6.2.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- 4034-PR-PRO-002 Sampling Protocol - the laboratory CN methods the analyser must track against

- KBK-MIR-MP-PRO-MET-SOP-0033 Penentuan Konsentrasi Sianida dengan Kolorimetri, received 11 Aug 2026

- SWI-PRO-MET-201-022 Online Analyser Verification Sampling, issued

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-207-012 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |