Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Fume Cupboard Operation, Scrubber Check and Cleaning

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-206-002 |
| --- | --- |
| Work instruction | SWI-PRO-MET-206-002 Fume Cupboard Operation, Scrubber Check and Cleaning |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Classification | Safety-critical |
| Register line | JSEA-PRO-MET-206-002 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task fume cupboard operation, scrubber check and cleaning in 206  Met Lab - Facility, Safety and Quality at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Operate the fume cupboards and prove the scrubber is working before any fume-generating work.

- The fume cupboard and its scrubber are the only thing between the technician and the fume from acid digestion, cyanide work and perchloric acid.

It is read with work instruction SWI-PRO-MET-206-002, which gives the step-by-step method, and with HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Confirm the extraction fan is running and the airflow indicator reads within its marked range before starting any work. | Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 2. Confirm the scrubber is running and its pH dosing is operating, and record the scrubber liquor pH. | Inhalation of acid fume or reagent vapour at the bench Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 3. Where scrubber liquor pH is out of range, stop - a scrubber outside its pH range is not scrubbing. | Inhalation of acid fume or reagent vapour at the bench Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 4. Work with the sash at the marked working height. Raising the sash breaks the containment. | Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 5. Keep apparatus at least 150 mm inside the sash line and never block the rear baffle. | Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 6. For perchloric acid work, use only the cupboard rated for it, and carry out the washdown routine immediately after every use. Perchlorate build-up in ductwork is explosive. | Inhalation of acid fume or reagent vapour at the bench Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 7. Wash down the working surface at the end of every session. | Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 8. Record the airflow reading, the scrubber pH and the washdown in the log. | Inhalation of acid fume or reagent vapour at the bench Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 2 | 4 | 8 MEDIUM | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Every use; scrubber checked at the start of every shift.

6.2  Job steps

1.  Confirm the extraction fan is running and the airflow indicator reads within its marked range before starting any work.

2.  Confirm the scrubber is running and its pH dosing is operating, and record the scrubber liquor pH.

3.  Where scrubber liquor pH is out of range, stop - a scrubber outside its pH range is not scrubbing.

4.  Work with the sash at the marked working height. Raising the sash breaks the containment.

5.  Keep apparatus at least 150 mm inside the sash line and never block the rear baffle.

6.  For perchloric acid work, use only the cupboard rated for it, and carry out the washdown routine immediately after every use. Perchlorate build-up in ductwork is explosive.

7.  Wash down the working surface at the end of every session.

8.  Record the airflow reading, the scrubber pH and the washdown in the log.

Equipment required:

- 2 x 2400 mm Safetyflow fume cupboards

- Model A crossflow scrubber with pH dosing

- Airflow indicator

- pH meter for scrubber liquor

- Washdown hose

- Log sheet

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-206-002 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-206-002.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-WET-SOP-054, -059 and ENV-SOP-061 fume cupboard and scrubber

- ILS quotation Q0023979 - 2 x 2400 mm Safetyflow fume cupboard rated for perchloric and HF, Model A crossflow scrubber, exhaust fans

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-206-002 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |