Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Met Lab Waste, Residue and Effluent Disposal

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-206-007 |
| --- | --- |
| Work instruction | SWI-PRO-MET-206-007 Met Lab Waste, Residue and Effluent Disposal |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Classification | Safety-critical |
| Register line | JSEA-PRO-MET-206-007 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task met lab waste, residue and effluent disposal in 206  Met Lab - Facility, Safety and Quality at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Dispose of laboratory waste, residues and effluent to the correct route.

- Cyanide-bearing laboratory effluent going to the wrong drain is an environmental incident and a reportable one.

It is read with work instruction SWI-PRO-MET-206-007, which gives the step-by-step method, and with HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Segregate at the bench. Cyanide-bearing, acidic, and general waste go to separate containers - never the same one. | Hydrogen cyanide gas released at the open sample point Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 2. NEVER put acidic waste into a cyanide-bearing waste container. That is an HCN generation event in a closed bin. | Hydrogen cyanide gas released at the open sample point Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 3. Neutralise acidic waste before it goes to any shared route, and confirm with pH paper. | Contact with sulphuric or hydrochloric acid - burn to skin or eyes Inhalation of acid fume or reagent vapour at the bench | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 4. Route cyanide-bearing effluent to the designated destination only. Confirm where that is before the first litre is generated. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 5. Label every waste container with contents, hazard and accumulation start date. | Manual handling injury from carrying sample buckets, containers and composite drums Cross-contamination from a container or cutter not cleaned between points | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 6. Log the waste in the register as it is generated, not from memory at the end of the week. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 7. Keep solid residues from testwork until the result is reported and accepted, then dispose. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 8. Arrange removal at the set interval and file the disposal documentation. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Every day, and at the end of every test.

6.2  Job steps

1.  Segregate at the bench. Cyanide-bearing, acidic, and general waste go to separate containers - never the same one.

2.  NEVER put acidic waste into a cyanide-bearing waste container. That is an HCN generation event in a closed bin.

3.  Neutralise acidic waste before it goes to any shared route, and confirm with pH paper.

4.  Route cyanide-bearing effluent to the designated destination only. Confirm where that is before the first litre is generated.

5.  Label every waste container with contents, hazard and accumulation start date.

6.  Log the waste in the register as it is generated, not from memory at the end of the week.

7.  Keep solid residues from testwork until the result is reported and accepted, then dispose.

8.  Arrange removal at the set interval and file the disposal documentation.

Equipment required:

- 750 L tipping bin

- 2 x 240 L Sulo bins

- 60 L drum

- Crossflow scrubber with pH control

- Waste labels

- Waste register

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-206-007 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-206-007.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-ENV-SOP-058 sludge transfer and ENV-SOP-106 container neutralisation

- ILS quotation Q0023979 - 750 L tipping bin, 2 x 240 L Sulo, 60 L drum, crossflow scrubber with pH control

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-206-007 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |