Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Drying Oven and Muffle Furnace Operation

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-206-010 |
| --- | --- |
| Work instruction | SWI-PRO-MET-206-010 Drying Oven and Muffle Furnace Operation |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Classification | Safety-critical |
| Register line | JSEA-PRO-MET-206-010 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task drying oven and muffle furnace operation in 206  Met Lab - Facility, Safety and Quality at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Operate the drying ovens, ashing furnace and muffle furnace safely and at the right temperature.

- Hot surfaces, hot glassware and a wrong set temperature that quietly ruins a batch of results.

It is read with work instruction SWI-PRO-MET-206-010, which gives the step-by-step method, and with HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Confirm the set temperature against the method before loading. A drying oven set to an ashing temperature destroys the sample and can start a fire. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Manual handling injury from carrying sample buckets, containers and composite drums Non-representative sample leading to a wrong process decision | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 2. Do not put volatile or flammable material in a drying oven. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 3. Load trays so air can circulate. An overloaded oven does not reach constant mass in the stated time. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Manual handling injury from carrying sample buckets, containers and composite drums | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 4. Use long tongs and heat-resistant gloves for everything going in or out. Set hot items on a heat mat, never on the bench. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 5. Cool samples in the desiccator before weighing, not on the open bench. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Inhalation of or contact with solid reagent - cyanide briquettes, lime or salt Non-representative sample leading to a wrong process decision | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Sampling of solid cyanide done only with the full cyanide PPE set including a gas mask with an A2B2E2K2P3 cartridge. Dust suppression and local extraction confirmed running. No dry sweeping. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 4 | 12 HIGH | Met Lab Technician |
| 6. For the ashing and muffle furnaces, confirm the extraction is running before the door is opened. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 7. Never open a muffle furnace at temperature with your face in front of the door. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. | 3 | 3 | 9 MEDIUM | Met Lab Technician |
| 8. Record the oven or furnace identity, set temperature and period against every batch, and log the daily temperature check. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Non-representative sample leading to a wrong process decision | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Every drying and ashing operation.

6.2  Job steps

1.  Confirm the set temperature against the method before loading. A drying oven set to an ashing temperature destroys the sample and can start a fire.

2.  Do not put volatile or flammable material in a drying oven.

3.  Load trays so air can circulate. An overloaded oven does not reach constant mass in the stated time.

4.  Use long tongs and heat-resistant gloves for everything going in or out. Set hot items on a heat mat, never on the bench.

5.  Cool samples in the desiccator before weighing, not on the open bench.

6.  For the ashing and muffle furnaces, confirm the extraction is running before the door is opened.

7.  Never open a muffle furnace at temperature with your face in front of the door.

8.  Record the oven or furnace identity, set temperature and period against every batch, and log the daily temperature check.

Equipment required:

- 2 x drying oven 560 L and 70 trays

- Ashing furnace 3.5 kW

- Muffle furnace SF-3-SD 3 L

- Long tongs

- Heat-resistant gloves

- Heat mats

- Desiccator

- Oven log

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-206-010 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-206-010.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-PRE-SOP-003 Pengeringan Sampel and WET-SOP-024 ashing furnace

- ILS quotation Q0023979 - 2 x drying oven 560 L with 70 trays, ashing furnace 3.5 kW, muffle furnace SF-3-SD

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-206-010 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |