Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Met Lab Emergency Response - Power Failure and Evacuation

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-206-011 |
| --- | --- |
| Work instruction | SWI-PRO-MET-206-011 Met Lab Emergency Response - Power Failure and Evacuation |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Classification | Safety-critical |
| Register line | JSEA-PRO-MET-206-011 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task met lab emergency response - power failure and evacuation in 206  Met Lab - Facility, Safety and Quality at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Respond to a power failure or an evacuation order in the laboratory.

- On a power failure the fume cupboards and the scrubber stop. Everything that was safe because it was being extracted is no longer safe.

It is read with work instruction SWI-PRO-MET-206-011, which gives the step-by-step method, and with HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. On loss of power, stop all work immediately. Do not continue by torchlight. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 2. Assume the fume cupboards and the scrubber have stopped. Anything generating fume is now venting into the room. | Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 3. Cap or cover open cyanide and acid containers if it can be done in seconds without leaning into a cupboard. If not, leave them. | Hydrogen cyanide gas released at the open sample point Inhalation of acid fume or reagent vapour at the bench Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 4. Switch off heating equipment - hotplates, ovens, furnaces - so they do not restart unattended when power returns. | Contact with hot solution or hot surface at the elution, stripping or regeneration circuits Contact with live electrical equipment at a sample pump, cutter or local control station | Confirm the circuit is not on a heating cycle before sampling. Heat-resistant gloves and face shield. Allow the sample line to flush and cool before collecting. Do not sample a column under pressure. No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 5. Leave the laboratory, close the door behind you and report to the muster point. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 6. Account for everyone who was in the laboratory. | Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 7. Do not re-enter until power is restored, extraction is confirmed running and the Laboratory Supervisor releases the room. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 8. On re-entry, check every instrument's state before restarting work, and discard any test that was interrupted rather than resuming it. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: On any power failure or evacuation; drill at the set interval.

6.2  Job steps

1.  On loss of power, stop all work immediately. Do not continue by torchlight.

2.  Assume the fume cupboards and the scrubber have stopped. Anything generating fume is now venting into the room.

3.  Cap or cover open cyanide and acid containers if it can be done in seconds without leaning into a cupboard. If not, leave them.

4.  Switch off heating equipment - hotplates, ovens, furnaces - so they do not restart unattended when power returns.

5.  Leave the laboratory, close the door behind you and report to the muster point.

6.  Account for everyone who was in the laboratory.

7.  Do not re-enter until power is restored, extraction is confirmed running and the Laboratory Supervisor releases the room.

8.  On re-entry, check every instrument's state before restarting work, and discard any test that was interrupted rather than resuming it.

Equipment required:

- Emergency lighting

- UPS-backed instruments

- Emergency contact list

- Muster point

- Personal HCN monitor

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-206-011 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-206-011.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- KBK-MIR-LAB-WET-SOP-050 Evacuation on sudden power failure

- ILS quotation Q0023979 - 2 x Riello UPS 2200VA

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-206-011 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |