Mt. Morgan JOB SAFETY AND ENVIRONMENTAL ANALYSIS

Met Lab Equipment Receipt, Installation and Commissioning Handover

Job Safety and Environmental Analysis  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | JSEA-PRO-MET-206-014 |
| --- | --- |
| Work instruction | SWI-PRO-MET-206-014 Met Lab Equipment Receipt, Installation and Commissioning Handover |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Classification | Operating |
| Register line | JSEA-PRO-MET-206-014 in HM-MMM-WHS-REG-JSEA-PRO-updated_10, tab JSEA_MET |
| Process owner | Heritage Mineral Operations · Business unit Mt. Morgan |
| Document owner | Process Operations Superintendent |
| Document approver | Process Manager |
| Revision | Rev 0 |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  PURPOSE / SCOPE

This job safety and environmental analysis covers the task met lab equipment receipt, installation and commissioning handover in 206  Met Lab - Facility, Safety and Quality at the Mt. Morgan processing plant. It identifies the hazards at each step of the task, the controls required, and the residual risk once those controls are in place.

- Receive, install and commission the laboratory equipment, and hand it over as fit for use.

- The laboratory is being bought as one package ex works. Nothing in it works until it is installed, connected, calibrated and proven.

It is read with work instruction SWI-PRO-MET-206-014, which gives the step-by-step method, and with HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality, the procedure this task sits under.

# 2  RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician | Only carry out this task where competent and authorised to do so. Wear the PPE listed in Section 5 without exception. Stop the task and report immediately if any control in Section 4 is not in place. Record the work accurately. Never invent or back-fill a result that was not obtained. |
| Supervisor / Coordinator | Ensure the safety of all personnel carrying out this task. Understand that much of the plant is automated and equipment may start or stop without warning. Confirm the task is authorised and that the area is fit before it starts. Ensure this JSEA is communicated to everyone doing the task and that the sign-on sheet is completed. |
| Manager / Superintendent | Ensure all personnel have received adequate training and are competent. Ensure this JSEA is reviewed at the interval in Section 10 and after any incident or change. Approve the risk ranking in Section 4. |

# 3  ABBREVIATIONS / DEFINITIONS

| Term | Meaning |
| --- | --- |
| AVR | Acidification, volatilisation and regeneration - the cyanide recovery step in Area 081 |
| CIL | Carbon in leach |
| CN Free | Free cyanide, determined by silver nitrate titration |
| DO | Dissolved oxygen |
| HCN | Hydrogen cyanide gas |
| JSEA | Job safety and environmental analysis - this document |
| P80 | The screen size through which 80 per cent of the material passes |
| PPE | Personal protective equipment |
| ReCYN | The resin-in-pulp cyanide and metal recovery process used in Areas 041, 051, 071 and 081 |
| SWI | Standard work instruction - the step-by-step document that sits under this JSEA |
| SWP | Standard work procedure - the circuit-level procedure this task sits under |
| TSF | Tailings storage facility |
| WAD CN | Weak acid dissociable cyanide |

# 4  POTENTIAL HAZARDS

| Scale | Values |
| --- | --- |
| Probability | 1 Rare 2 Unlikely 3 Possible 4 Likely 5 Almost certain |
| Consequence | 1 Insignificant 2 Minor 3 Moderate 4 Major 5 Catastrophic |
| Risk rank | Probability x Consequence. 1-4 Low 5-9 Medium 10-14 High 15-25 Extreme. Scored after the controls in the same row are in place. |

| Job step | Hazards identified | Controls required | P | C | Risk | Actioned by |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Check every delivery against quotation Q0023979 line by line. The quote is ex works, so shortages and transit damage are found here or not at all. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 2. Confirm the services each item needs are in place before installation - power, water, drainage, compressed air, extraction. | Contact with live electrical equipment at a sample pump, cutter or local control station Release from a pressurised sample line, hose or fitting | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 3. Note the items that need work not in the vendor scope: the dust collector pulse-system wiring, its isolator, the GPO and the compressed air connection. | Manual handling injury from carrying sample buckets, containers and composite drums Release from a pressurised sample line, hose or fitting Non-representative sample leading to a wrong process decision | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 4 | 12 HIGH | Met Lab Technician |
| 4. Register the 600 L air receiver as a pressure vessel and arrange its statutory inspection before it is put into service. This is a legal requirement in Queensland, not a commissioning nicety. | Release from a pressurised sample line, hose or fitting | Confirm the sample line is not blocked before opening. Open slowly and stand clear of the discharge path. Report any weeping fitting rather than tightening it under pressure. | 2 | 4 | 8 MEDIUM | Met Lab Technician |
| 5. Install, level and connect each instrument per its manual. | Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. | 2 | 5 | 10 HIGH | Met Lab Technician |
| 6. Calibrate or verify every instrument and balance before it is used for any result, and file the certificates. | Contact with live electrical equipment at a sample pump, cutter or local control station Non-representative sample leading to a wrong process decision | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 7. Run a commissioning check on each instrument against a known standard and record the result. | Contact with live electrical equipment at a sample pump, cutter or local control station Non-representative sample leading to a wrong process decision | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 5 | 15 EXTREME | Met Lab Technician |
| 8. Enter every item in the asset register with serial number, location, calibration interval and PM requirement. | Manual handling injury from carrying sample buckets, containers and composite drums | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. | 3 | 2 | 6 MEDIUM | Met Lab Technician |
| 9. Hand over formally, item by item, and record what is accepted and what is outstanding. | Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. | 3 | 3 | 9 MEDIUM | Met Lab Technician |

# 5  PPE

Standard PPE is worn for all normal operations in the Mt. Morgan gold processing plant. The following is required for this task.

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 6  PROCEDURE

6.1  Introduction

Warning: many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times.

Points and equipment covered by this task:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Once, at laboratory establishment; then on every new instrument.

6.2  Job steps

1.  Check every delivery against quotation Q0023979 line by line. The quote is ex works, so shortages and transit damage are found here or not at all.

2.  Confirm the services each item needs are in place before installation - power, water, drainage, compressed air, extraction.

3.  Note the items that need work not in the vendor scope: the dust collector pulse-system wiring, its isolator, the GPO and the compressed air connection.

4.  Register the 600 L air receiver as a pressure vessel and arrange its statutory inspection before it is put into service. This is a legal requirement in Queensland, not a commissioning nicety.

5.  Install, level and connect each instrument per its manual.

6.  Calibrate or verify every instrument and balance before it is used for any result, and file the certificates.

7.  Run a commissioning check on each instrument against a known standard and record the result.

8.  Enter every item in the asset register with serial number, location, calibration interval and PM requirement.

9.  Hand over formally, item by item, and record what is accepted and what is outstanding.

Equipment required:

- ILS quotation Q0023979 as the asset list

- Installation drawings and services schedule

- Calibration certificates

- Commissioning checklists

- Asset register

# 7  FLOWCHART

Not reproduced in this document. The step sequence is in work instruction SWI-PRO-MET-206-014 Section 6, which is the controlled version of the method.

# 8  TRAINING

Every person carrying out this task has been through this JSEA at a communication session recorded on the sign-on sheet at the end of this document, and is trained and signed off against work instruction SWI-PRO-MET-206-014.

- Site induction, current.

- Cyanide awareness and gas detection competency where the area requires it.

- Working at height competency where the task requires it.

- Competency in the determinations and calculations carried out in this task.

# 9  REFERENCE

- ILS quotation Q0023979 - the asset list, ex works, 12 to 18 week lead time, with 10 per cent of value payable against commissioning

- The quotation's stated validity expired 16 February 2026 - it is a quotation, not a purchase order

# 10  REVIEW

This JSEA is reviewed every 12 months, and immediately after any incident, near miss, plant change or change to the schedule affecting this task.

# JSEA COMMUNICATION AND AWARENESS

| Presented by | ________________________ |
| --- | --- |
| Date | ________________________ |
| JSEA reference | JSEA-PRO-MET-206-014 |

| Name | Position | Department | Signature |
| --- | --- | --- | --- |